Table S3.  Evaluation of the detection efficiency for S. pneumoniae and negative in clinical samples

|  | S. pneumoniae | S. pyogenes | Negative |
| --- | --- | --- | --- |
|  | Ref© + | Ref© - | Ref© + | Ref© - | Ref © + | Ref © - |
| RPM v.1 + | 36 | 2 | 13 | 0 | 52 | 7 |
| RPM v.1 - | 5 | 381 | 0 | 411 | 0 | 365 |
| Sensitivity | 88% | 100 % | 100% |
| Specificity | 99.5% | 100 % | 98% |
| Overall agreement | 98% | 100 % | 98% |